Table 1 Percentage of total allowable error (TEa), bias, CV for Phase 1,2,3 at level (L1) and sigma metric of analytes

| Analytes | Abbreviation | Unit | TEa(%) | Phase 1 % Bias | Phase 2 % Bias | Phase 3 % Bias | Phase 1 CV% | Phase 2 CV% | Phase 3 CV% | Phase 1 Sigma (L1) | Phase 2 Sigma (L1) | Phase 3 Sigma (L1) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Albumin | ALB | g/dL | 10 | 1.1 | 2.1 | 0.9 | 2.3 | 3.1 | 2.4 | 3.9 | 2.5 | 3.8 |
| Alkaline phosphatase | ALKP | U/L | 30 | 3.2 | 2.6 | 6.9 | 2.1 | 3.7 | 2.8 | 12.8 | 7.4 | 8.3 |
| Alanine aminotransferase | ALT | U/L | 20 | 7.2 | 4.9 | 4.3 | 10.3 | 10.7 | 8.7 | 1.2 | 1.4 | 1.8 |
| Amylase | AMY | U/L | 30 | 1.4 | 4.9 | 3.7 | 3.9 | 4.4 | 5.2 | 7.3 | 5.7 | 5.1 |
| Aspartate aminotransferase | AST | U/L | 20 | 1.1 | 1.3 | 1.3 | 2.1 | 2.3 | 2.2 | 9.0 | 8.1 | 8.5 |
| Bicarbonate | CO2 | mmol/L | 10 | 5.6 | 8.5 | 2.0 | 3.4 | 4.7 | 5.2 | 1.3 | 0.3 | 1.5 |
| Bilirubin, total | TBIL | mg/dL | 20 | 4.8 | 1.0 | 1.4 | 4.8 | 6.4 | 6.0 | 3.2 | 3.0 | 3.1 |
| Calcium | Ca | mg/dL | 11 | 1.7 | 3.1 | 1.1 | 1.3 | 0.9 | 1.1 | 7.2 | 8.8 | 9.0 |
| Cholesterol, total | CHOL | mg/dL | 10 | 2.2 | 0.3 | 1.9 | 1.8 | 2.4 | 1.9 | 4.3 | 4.0 | 4.3 |
| Creatine kinase | CK | U/L | 30 | 3.5 | 4.6 | 2.1 | 2.7 | 6.9 | 3.7 | 9.8 | 4.0 | 7.5 |
| Chloride | Cl- | mmol/L | 5 | 1.1 | 1.1 | 0.4 | 1.6 | 1.8 | 1.1 | 2.4 | 2.2 | 4.2 |
| Creatinine | CREAT | mg/dL | 15 | 4.1 | 3.1 | 2.3 | 1.3 | 1.7 | 1.9 | 8.4 | 7.0 | 6.7 |
| HDL cholesterol | HDLC | mg/dL | 30 | 7.5 | 7.6 | 5.2 | 2.7 | 3.8 | 3.9 | 8.3 | 5.9 | 6.4 |
| LDL cholesterol | LDLC | mg/dL | 22 | 4.1 | 10.1 | 2.7 | 3.0 | 1.9 | 2.8 | 6.0 | 6.3 | 6.9 |
| Iron, total | Fe | ug/dL | 20 | 7.2 | 5.7 | 6.3 | 4.7 | 3.5 | 2.6 | 2.7 | 4.1 | 5.3 |
| gamma GGT | GGT | U/L | 22.2 | 6.5 | 0.8 | 1.0 | 2.0 | 2.5 | 1.6 | 7.9 | 8.6 | 13.3 |
| Glucose | GLU | mg/dL | 10 | 1.7 | 0.6 | 1.3 | 1.4 | 1.5 | 1.3 | 5.9 | 6.3 | 6.7 |
| Potassium | K+ | mmol/L | 5 | 0.3 | 1.3 | 0.8 | 1.7 | 2.2 | 1.2 | 2.8 | 1.7 | 3.5 |
| Lactate dehydrogenase | LDH | U/L | 20 | 2.2 | 4.0 | 3.2 | 3.5 | 3.4 | 5.0 | 5.1 | 4.7 | 3.4 |
| Lipase | LIPA | U/L | 20 | 1.5 | 1.3 | 3.7 | 1.9 | 2.7 | 1.6 | 9.7 | 6.7 | 10.2 |
| Magnesium | Mg | mg/dL | 25 | 2.2 | 4.4 | 3.6 | 1.8 | 2.2 | 2.5 | 12.7 | 9.4 | 8.6 |
| Sodium | Na+ | mmol/L | 5 | 0.4 | 1.5 | 0.9 | 1.3 | 2.1 | 0.9 | 3.5 | 1.7 | 4.6 |
| Phosphorous | PHOS | mg/dL | 10 | 2.8 | 3.1 | 2.5 | 1.4 | 2.1 | 1.9 | 5.1 | 3.3 | 3.9 |
| Total protein | TP | mg/dL | 10 | 2.8 | 3.7 | 2.9 | 1.2 | 1.9 | 1.2 | 6.0 | 3.3 | 5.9 |
| Triglycerides | TRIG | mg/dL | 25 | 2.6 | 0.6 | 2.0 | 1.7 | 1.7 | 1.5 | 13.2 | 14.4 | 15.3 |
| Urea | UREA | mg/dL | 19.2 | 1.1 | 3.2 | 0.3 | 1.2 | 3.8 | 2.0 | 15.1 | 4.2 | 9.5 |
| Uric acid | UA | mg/dL | 17 | 1.1 | 1.1 | 2.6 | 1.2 | 1.1 | 1.4 | 13.8 | 14.5 | 10.3 |

Table 2 Percentage of total allowable error (TEa), bias, CV for Phase 1,2,3 at level (L2) and sigma metric  of analytes

| Analytes | Abbreviation | Unit | TEa(%) | Phase 1 % Bias | Phase 2 % Bias | Phase 3 % Bias | Phase 1 CV% | Phase 2 CV% | Phase 3 CV% | Phase 1 Sigma (L2) | Phase 2 Sigma (L2) | Phase 3 Sigma (L2) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Albumin | ALB | g/dL | 10 | 1.1 | 2.1 | 0.9 | 2.5 | 2.4 | 2.7 | 3.6 | 3.3 | 3.4 |  |
| Alkaline phosphatase | ALKP | U/L | 30 | 3.2 | 2.6 | 6.9 | 2.1 | 3.2 | 2.4 | 12.8 | 8.6 | 9.6 |  |
| Alanine aminotransferase | ALT | U/L | 20 | 7.2 | 4.9 | 4.3 | 2.3 | 3.6 | 2.9 | 5.6 | 4.2 | 5.1 |  |
| Amylase | AMY | U/L | 30 | 1.4 | 4.9 | 3.7 | 1.7 | 2.1 | 2.9 | 16.8 | 12.0 | 9.1 |  |
| Aspartate aminotransferase | AST | U/L | 20 | 1.1 | 1.3 | 1.3 | 1.4 | 2.1 | 1.7 | 13.5 | 8.9 | 11.0 |  |
| Bicarbonate | CO2 | mmol/L | 10 | 5.6 | 8.5 | 2.0 | 3.9 | 7.2 | 6.0 | 1.1 | 0.2 | 1.3 |  |
| Bilirubin, total | TBIL | mg/dL | 20 | 4.8 | 1.0 | 1.4 | 2.7 | 4.3 | 3.7 | 5.6 | 4.4 | 5.0 |  |
| Calcium | Ca | mg/dL | 11 | 1.7 | 3.1 | 1.1 | 1.1 | 1.2 | 1.4 | 8.5 | 6.6 | 7.1 |  |
| Cholesterol, total | CHOL | mg/dL | 10 | 2.2 | 0.3 | 1.9 | 1.3 | 1.9 | 1.6 | 6.0 | 5.1 | 5.1 |  |
| Creatine kinase | CK | U/L | 30 | 3.5 | 4.6 | 2.1 | 2.4 | 5.0 | 3.1 | 11.0 | 5.1 | 9.0 |  |
| Chloride | Cl- | mmol/L | 5 | 1.1 | 1.1 | 0.4 | 1.6 | 1.6 | 1.1 | 2.4 | 2.4 | 4.2 |  |
| Creatinine | CREAT | mg/dL | 15 | 4.1 | 3.1 | 2.3 | 1.5 | 2.1 | 1.5 | 7.3 | 5.7 | 8.5 |  |
| HDL cholesterol | HDLC | mg/dL | 30 | 7.5 | 7.6 | 5.2 | 2.1 | 2.9 | 4.4 | 10.7 | 7.7 | 7.6 |  |
| LDL cholesterol | LDLC | mg/dL | 22 | 4.1 | 10.1 | 2.7 | 3.5 | 3.4 | 2.9 | 5.1 | 3.5 | 6.7 |  |
| Iron, total | Fe | ug/dL | 20 | 7.2 | 5.7 | 6.3 | 6.2 | 5.1 | 4.2 | 2.1 | 2.8 | 3.3 |  |
| gamma GGT | GGT | U/L | 22.2 | 6.5 | 0.8 | 1.0 | 1.4 | 1.6 | 1.6 | 11.2 | 13.4 | 13.3 |  |
| Glucose | GLU | mg/dL | 10 | 1.7 | 0.6 | 1.3 | 1.1 | 1.4 | 1.2 | 7.5 | 6.7 | 7.3 |  |
| Potassium | K+ | mmol/L | 5 | 0.3 | 1.3 | 0.8 | 2.1 | 2.7 | 1.0 | 2.2 | 1.4 | 5.0 |  |
| Lactate dehydrogenase | LDH | U/L | 20 | 2.2 | 4.0 | 1.1 | 2.1 | 2.6 | 3.1 | 8.5 | 6.2 | 6.1 |  |
| Lipase | LIPA | U/L | 20 | 1.5 | 1.3 | 3.7 | 1.7 | 1.9 | 1.8 | 10.9 | 9.8 | 9.1 |  |
| Magnesium | Mg | mg/dL | 25 | 2.2 | 4.4 | 3.6 | 1.4 | 2.0 | 1.6 | 16.3 | 10.3 | 13.4 |  |
| Sodium | Na+ | mmol/L | 5 | 0.4 | 1.5 | 0.9 | 1.5 | 2.0 | 0.7 | 3.1 | 1.8 | 5.9 |  |
| Phosphorous | PHOS | mg/dL | 10 | 2.8 | 3.1 | 1.8 | 1.1 | 1.4 | 1.7 | 6.5 | 4.9 | 4.4 |  |
| Total protein | TP | mg/dL | 10 | 2.8 | 3.7 | 2.9 | 1.1 | 2.1 | 1.1 | 6.5 | 3.0 | 6.5 |  |
| Triglycerides | TRIG | mg/dL | 25 | 2.6 | 0.6 | 2.0 | 1.4 | 2.4 | 2.1 | 16.0 | 10.2 | 11.0 |  |
| Urea | UREA | mg/dL | 19.2 | 1.1 | 3.2 | 0.3 | 1.2 | 3.0 | 2.0 | 15.1 | 5.3 | 9.5 |  |
| Uric acid | UA | mg/dL | 17 | 1.1 | 1.1 | 2.6 | 1.4 | 1.4 | 1.8 | 11.4 | 11.4 | 8.0 |  |

Table 3 The criteria we use for interpreting QGI when test applications fall short of six sigma quality is as follows

| QGI | Problem |
| --- | --- |
| <0.8 | Imprecision |
| 0.8 - 1.2 | Imprecision & Inaccuracy |
| >1.2 | Inaccuracy |

Table 4 shows list of analytes performed low during Phase 2 assessed for Sigma QGI for Accuracy and Precision problem.

| Analytes | Qc Levels | %Bias | %CV | Sigma | QGI | Problem |
| --- | --- | --- | --- | --- | --- | --- |
| Iron | Level 1 | 5.7 | 3.5 | 4.1 | 1.1 | Imprecision & Inaccuracy |
|  | Level 2 | 5.7 | 5.1 | 2.8 | 0.7 | Imprecision |
| ALB | Level 1 | 2.1 | 3.1 | 2.5 | 0.5 | Imprecision |
|  | Level 2 | 2.1 | 2.4 | 3.3 | 0.6 | Imprecision |
| ALT | Level 1 | 4.9 | 10.7 | 1.4 | 0.3 | Imprecision |
|  | Level 2 | 4.9 | 3.6 | 4.2 | 0.9 | Imprecision & Inaccuracy |
| bicarbonate | Level 1 | 8.5 | 4.7 | 0.3 | 1.2 | Imprecision & Inaccuracy |
|  | Level 2 | 8.5 | 7.2 | 0.2 | 0.8 | Imprecision & Inaccuracy |
| sodium | Level 1 | 1.5 | 2.1 | 1.7 | 0.5 | Imprecision |
|  | Level 2 | 1.5 | 2 | 1.8 | 0.5 | Imprecision |
| potassium | Level 1 | 1.3 | 2.2 | 1.7 | 0.4 | Imprecision |
|  | Level 2 | 1.3 | 2.7 | 1.4 | 0.3 | Imprecision |
| Chloride | Level 1 | 1.1 | 1.8 | 2.2 | 0.4 | Imprecision |
|  | Level 2 | 1.1 | 1.6 | 2.4 | 0.5 | Imprecision |

Table 5 Shows ALT, Bicarbonate, Phase 3 six sigma, QGI value, %CV comparison with Manufacturer claim

| Analytes | QC Levels | %Bias | %CV | Sigma | QGI | Problem | Manufacturer's Claim CV% |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ALT | Level 1 | 4.3 | 8.7 | 1.8 | 0.3 | Imprecision | 8.5 |
| Bicarbonate | Level 1 | 2 | 5.2 | 1.5 | 0.3 | Imprecision | 4.7 |
|  | Level 2 | 2 | 6 | 1.3 | 0.2 | Imprecision | 3.9 |